## Slide 1: BigMeasure Trigger Configurations
K. Fine
J. Eagan
03 Dec 2009

## Slide 2: Purpose
Setup of multiple DAQs for use with BigMeasure program.
Describe connections between NI DAQs when multiple DAQs are present.

## Slide 3: Important Points
Justin E. and Kevin F. have tested these modes on NI-6259 and NI-6225 with one slave.
Analog Channel trigger not available on NI-6255.
All wiring between DAQs is same between trig modes. Only difference is external trigger for Master.
Testing showed that connections between DAQs should be short and twisted pair or they will not trigger reliably.

## Slide 4: Immediate Trigger DAQ is triggered when GO button of BigMeasure pressed
Master Properties:
TriggerType = Immediate
ExternalTriggerDriveLine = PFI10
ClockSource = (Internal)
ExternalSampleClockDriveLine = PFI11
ExternalScanClockDriveLine = PFI12
Slave Properties:
TriggerType = HwDigital
HwDigitalTriggerSource = PFI10
TriggerCondition = Negative Edge
ClockSource = ExternalSampleAndScanCtrl
ExternalSampleClockSource = PFI11
ExternalScanClockSource = PFI12
Master
Slave 1
Slave 2
Slave 3
PFI10
Trig In
PFI10
Trig Out
PFI10
Trig In
PFI10
Trig In
PFI12
Scan Clock Out
PFI12
Scan Clock In
PFI12
Scan Clock In
PFI12
Scan Clock In
PFI11
Sample Clock Out
PFI11
Sample Clock In
PFI11
Sample Clock In
PFI11
Sample Clock In

## Slide 5: External Digital Trigger DAQ is triggered by external TTL
Master
Slave 1
Slave 2
Slave 3
Master Properties:
TriggerType = HwDigital
TriggerCondition = Negative Edge
or Positive Edge
HwDigitalTriggerSource = PFI0-9 (user selectable)
ExternalTriggerDriveLine = PFI10
ClockSource = (Internal)
ExternalSampleClockDriveLine = PFI11
ExternalScanClockDriveLine = PFI12
Slave Properties:
TriggerType = HwDigital
HwDigitalTriggerSource = PFI10
TriggerCondition = Negative Edge
ClockSource = ExternalSampleAndScanCtrl
ExternalSampleClockSource = PFI11
ExternalScanClockSource = PFI12
PFI10
Trig In
External
Trigger
PFI10
Trig Out
PFI0-9
Trig In
PFI10
Trig In
PFI10
Trig In
PFI12
Scan Clock Out
PFI12
Scan Clock In
PFI12
Scan Clock In
PFI12
Scan Clock In
PFI11
Sample Clock Out
PFI11
Sample Clock In
PFI11
Sample Clock In
PFI11
Sample Clock In

## Slide 6: Analog Channel Trigger DAQ is triggered by first input channel (not available on NI USB-6225)
Master Properties:
TriggerType = HwAnalogChannel
TriggerCondition = AboveHighLevel
or BelowLowLevel
TriggerConditionValue = Trig Volts
ExternalTriggerDriveLine = PFI10
ClockSource = (Internal)
ExternalSampleClockDriveLine = PFI11
ExternalScanClockDriveLine = PFI12
Slave Properties:
TriggerType = HwDigital
HwDigitalTriggerSource = PFI10
TriggerCondition = Negative Edge
ClockSource = ExternalSampleAndScanCtrl
ExternalSampleClockSource = PFI11
ExternalScanClockSource = PFI12
Master
Slave 1
Slave 2
Slave 3
PFI10
Trig In
PFI10
Trig Out
PFI10
Trig In
PFI10
Trig In
PFI12
Scan Clock Out
PFI12
Scan Clock In
PFI12
Scan Clock In
PFI12
Scan Clock In
PFI11
Sample Clock Out
PFI11
Sample Clock In
PFI11
Sample Clock In
PFI11
Sample Clock In